Quality Assurance and Quality Control in the EAL

The EAL uses an appropriate Quality Assurance and Quality Control program (QA/QC) to ensure precision and accuracy when analyzing all samples. We use the QA/QC methods outlined in Soil, Plant and Water Reference Methods for the Western Region, 4th edition. A copy can be accessed on the EAL website. (https://pws.byu.edu/eal/resources)

Blanks: Blanks are control samples added to every batch of samples to be analyzed. The purpose is to act as a control to determine interference or contamination that may artificially affect the results of the actual samples. Blanks contain all components except for the specific substance (analyte) being measured. There are two types of blanks—Batch and Process Blanks. A Process Blank goes through most all of the process as samples being analyzed go through but do not contain any soil, plant material, or other materials being tested. Because we add nothing to this blank it should read near “zero” concentration and, thus, can be used to know if there was contamination error. A Batch or Calibration Blank enters the process during the actual instrument analysis and is used to identify and correct for any background signal used in calibration of the instrument and to monitor it.

Positive Control: The EAL uses two types of positive controls, external and internal checks. External checks are samples from an outside organization. The EAL participates in the North American Proficiency Testing program (NAPT) and receives samples from them twice a year. The NAPT samples are known as “blind” samples as we know they are QA/QC, but we do not know their values at the time of initial testing. After we analyze the samples, our results are graded by comparing to the Mean Absolute Deviation (MAD). Results outside +4.0 MAD are considered extreme outliers. Results between +2.5 and 4.0 MAD are deemed a “warning”.

Internal checks are similar to blanks as they are analyzed alongside every batch of samples. A large volume of dried material (e.g., soil, plant tissue, etc.) is prepared for analysis by fine grinding and homogenization. Liquid checks (e.g., irrigation water samples, etc.) are also prepared by stabilizing with acid and microbial suppressant. The internal checks are analyzed multiple times to develop statistical measures of precision. The previously used internal check and/or an external check (e.g., NAPT) are also analyzed with it initially to verify accuracy with a high degree of certainty. Once developed, an internal check sample is included in each batch of regular samples to be used to access accuracy of the batch. If the values for these internal check samples are abnormal some or all of the batch may need to be redone (considering all of the QA/QC results in the batch by the QA/QC officer).

Duplicates: Every batch of samples analyzed includes a duplicate (a sample we analyze two times). Every 15-20 samples include at least one duplicate sample. If a duplicate’s results are different by >10%, we know there was instrument or human error and can correct as needed. Occasionally instruments will “drift”. This and other errors can be identified by adding random duplicates.

Calibrations: All instruments in the EAL are calibrated regularly, according to the needs of each instrument. Calibrating regularly ensures the instrument is performing as it should and gives us accurate results. Calibration is done with “standards”, or solutions which have a known quantity. Instruments are calibrated to R2 >0.999, which indicates a high degree of accuracy.

Instrument Checks: We check for instrument drift by performing instrument checks. Checks are ran at the beginning, end, and after every 15 samples of a batch. These checks are standard analyzations, like positive controls. If the result of a check is off, the instrument is recalibrated before continuing.

Lab Quality Assurance/Quality Control (QA/QC) Officer: The EAL has a designated QA/QC officer who trains all technicians/analysts on QA/QC and reviews all results before they are sent. The current QA officer is Dr. Erin Jones, our lab manager. If any results fail to pass standards, Dr. Jones diagnoses the error and requires the samples to be re-analyzed.

Re-Analysis/Re-Running Samples: If after receiving results a customer suspects there is an error, they may request the EAL to re-analyze their samples. If after the second analysis there is a discrepancy of at least 10% between results, the EAL will not charge the customer for the re-analysis.